## Pass 2 Atomic and tagging
Important notice: This document is an attempt to analyse, review and clarify the 'RMSV001' document published by the Medicines and Healthcare products Regulatory Agency (MHRA). In all cases readers must refer to that document as the primary source of truth regarding the requirement. This document has not been reviewed or approved by the MHRA and as such is provided on a 'best-efforts' basis without any warranty of quality, accuracy or completeness.
Rapidly Manufactured Ventilator System
(RMVS) Document RMVS001 - Specification Issued by MHRA
Introduction
This is a specification of the minimally (and some preferred options) clinically acceptable ventilator to be used in UK hospitals during the current COVID-19 pandemic caused by SARS-CoV-2 virus. It sets out the clinical requirements based on the consensus of what is ‘minimally acceptable’ performance in the opinion of the anaesthesia and intensive care medicine professionals and medical device regulators. It is for devices, which are most likely to confer therapeutic benefit on a patient suffering with ARDS caused by SARS-CoV-2, used in the initial care of patients requiring urgent ventilation. A ventilator with lower specifications than this is likely to provide no clinical benefit and might lead to increased harm, which would be unacceptable for clinicians. |  |  | Comment | Requirement Priority | Comment
Intensive care medicine is a whole system of care and ventilators cannot be safely used on any patient without trained staff and other equipment and medicines. Where these impinge on the specification they are mentioned below. |  |  | Comment
It is proposed these ventilators would be for short-term stabilisation for a few hours, but this may be extended up to 1-day use for a patient in extremis as the bare minimum function. Ideally it would also be able to function as a broader function ventilator which could support a patient through a number of days, when more advanced ventilatory support becomes necessary. | 1 | The system shall enable short-term stablisation for up to [Duration TBD ] Hours | REQT | Mandatory
| 2 | The system shall function as a broader-function ventilator for up to [Duration TBD Assume 5] Days | REQT | Desirable
Ventilation
1. At least 1, optionally 2 modes of ventilation |  |  | REQT
a. Must have mandatory ventilation. | 3 | The system shall provide ventilation | REQT | Mandatory | Unclear. Is this saying that there is a type of ventilation called "mandatory ventillation" or is it saying that it is mandatory (in requirement terms) that the device provides one of the following forms of ventilation?
b. Mandatory mode must be either (ideally) Pressure Regulated Volume Control, or a pressure controlled ventilation (PCV) or a volume controlled ventilation (VCV). | 4 | The system shall provide ventilation using Pressure Regulated Volume Control (PRVC) | REQT | Mandatory
| 5 | The system shall provide ventilation using Pressure Controlled Ventilation (PCV) | REQT | Optional
|  | Pressure Control Ventilation (PVC) is a mode in which the user has to provide the adaptive control to achieve tidal volume | Comment
| 6 | The system shall provide ventilation using Volume Controlled Ventilation (VCV). | REQT | Optional
c. PRVC/Pressure Controlled - a set pressure is delivered for the period of inspiration and the volume achieved is measured and displayed. Ideally PRVC, an adaptive mode where the tidal volume is set and the lowest possible pressure is delivered to achieve this volume. Pressure Control Ventilation where the user has to provide the adaptive control to achieve tidal volume is only acceptable if the tidal volume delivered is clearly displayed and the user can set patient specific upper and lower tidal volume alarms to alert to the need to adjust the pressure. | 7 | If a PRVC mode is provided then the system shall deliver a set pressure for the period of inspiration | REQT | Optional
| 8 | If a PRVC mode is provided then the system shall display the volume achieved. | REQT | Optional | Unclear. What does 'Volume achieved' mean? Is this the same as "Tidal Volume Achieved"? Does this mean the total volume of air inhaled by the patient?
| 9 | If a PRVC mode is provided then the system shall additionally provide an 'adaptive' mode in which the tidal-volume is set and the lowest possible pressure is delivered to achieve this volume. | REQT | Optional
| 10 | If a Pressure Controlled Ventillation (PCV) mode is provided then the system shall [ clearly ] display the tidal volume | REQT | Mandatory
| 11 | If a Pressure Controlled Ventillation (PCV) mode is provided then the system shall enable a user to set both upper and lower tidal volume alert levels | REQT | Mandatory
| 12 | If a Pressure Controlled Ventillation (PCV) mode is provided then the system shall shall raise an alarm if either the upper or lower tidal alarm levels are not reached | REQT | Mandatory
d. Volume Control – the user sets a tidal volume and respiratory rate. The tidal volume is delivered during the inspiratory period. Acceptable only if additional pressure limiting controls are available, see Inspiratory Pressure section. |  | Volume Controlled Ventillation is a mode in which the user sets a tidal volume and respiratory rate. The tidal volume is delivered during the inspiratory period. | Comment
| 13 | If the system has a Volume Controlled Ventillation mode then the system shall include pressure limiting controls (See Inspiratory Pressure section) | REQT | Mandatory
e. Optional pressure support mode for those patients breathing to some extent themselves, e.g. BIPAP or SIMV-PC. The user sets an inspiratory pressure and an expiratory pressure. The ventilator can sense when a patient starts to breathe in and apply the inspiratory pressure, then sense when the patient starts to breathe out and apply the expiratory pressure (this pressure is still positive but lower than the inspiratory pressure). |  | Pressure support mode' (e.g. BIPAP or SIMV-PC) is used for those patients breathing to some extent themselves, The user sets an inspiratory pressure and an expiratory pressure. The ventilator can sense when a patient starts to breathe in and apply the inspiratory pressure, then sense when the patient starts to breathe out and apply the expiratory pressure (this pressure is still positive but lower than the inspiratory pressure). | Comment
| 14 | The system shall provide a BIPAP (pressure support) mode | REQT | Optional
| 15 | The system shall provide a SIMV-PC (pressure support) mode | REQT | Optional
2. If the patient stops breathing in a spontaneous mode it must failsafe automatically onto mandatory ventilation. |  | If the system provides a pressure support mode then the system must automatically detect when the patient stops breathing and enter a 'failsafe' mode in which the system provides [ mandatory ] ventilation. |  |  | Unclear: What is "Spontaneous Mode" this has not been defined in the previous text. Assume they are still talking about "Pressure Support Mode" There is another use of "Mandatory ventilation" here. Assume that this means a mode in which the system forces air into the patient.
3. Inspiratory airway pressure, the higher pressure setting that is applied to make the patient breathe in: |  | Inspiratory airway pressure', is the higher pressure that is applied to make the patient breathe in | Comment
a. Plateau pressure should be adjusted to achieve volume and be limited to 35 cmH2O by default. It is acceptable if an option to increase this to 70 cmH2O in exceptional circumstances is provided | 16 | The system shall enable a plateu pressure for the inspiratory airway pressure to be set by the user. | REQT | Optional
| 17 | The system shall provide a 'default' mode in which the plateu pressure of the inspiratory airway pressure is limited to 35 cmH2O | REQT | Mandatory
| 18 | The system shall enable a user to override the default maximum setting for the plateu pressure of the inspiratory airway pressure. | REQT | Optional
| 19 | If the system provides an override for the default maximum setting for the plateau pressure of the inspiratory airway system then in any case the absolute maximum pressure shall be limited to 70 cmH2O | REQT | Mandatory
b. Peak pressure should be no more than 2 cmH2O greater than plateau pressure. | 20 | The system shall ensure that the peak pressure is no greater than 2 cmH20 above the [ user defined ] plateu pressure. | REQT | Mandatory
c. If volume control ventilation is used, the user must be able to set inspiratory airway pressure limit in the range at least 15 – 40 cmH2O in at least increments of 5 cmH2O. | 21 | If the system provides Volume Control Ventilation then the system shall enable the user to set the airway pressure within the range 15 - 40 cmH2O | REQT | Mandatory | Is this saying that a wider range would be acceptable?
| 22 | If the system provides Volume Control Ventilation then the system shall enable the user to set the airway pressure in increments of 5cmH2O | REQT | Mandatory
d. There must be a mechanical failsafe valve that opens at 80 cmH2O. | 23 | The system shall provide a mechanical failsafe valve that opens at 80 cmH2O | REQT | Mandatory
4. Positive End Expiratory Pressure (PEEP). The pressure maintained in the breathing system during expiration. |  | 'Positive End Expiratory Pressure' (PEEP) is the pressure maintained in the breathing system during expiration. | Comment
a. Range 5-20 cm H2O adjustable in 5 cmH2O increments. | 24 | The system shall provide a PEEP. | REQT | Mandatory
| 25 | The system shall enable a user t adjust the PEEP in the range 5 to 20 cmH20 | REQT | Mandatory
| 26 | The system shall enable a user t adjust the PEEP in increments of 5cm H2O | REQT | Mandatory
b. Patient breathing system must remain pressurised to at least the PEEP level setting at all times. | 27 | The system shall ensure that the patient breathing system is pressurised to the defined PEEP level at all times | REQT | Mandatory | Requirement says "at least to the PEEP level" - this can't be correct. Surelly AT the PEEP level +/- some error bar?
5. Inspiratory:Expiratory ratio (I:E). The proportion of each breathing cycle that is spent breathing in compared to breathing out. |  | 'Inspiratory:Expiratory ratio' (I:E) is the proportion of each breathing cycle that is spent breathing in compared to breathing out. | Comment
a. 1:2.0 (i.e. expiration lasts twice as long as inspiration). | 28 | The system shall provide an Inspiratory Expiration ratio of 1:2.0 (I.e. expiration shall take twice as lons as inspiration) | REQT | Mandatory | Why does the '1' have no decimal but the second number is shown as '2.0'? Assume typo rather than an indicator of accuracy.
b. Optionally adjustable in the range 1:1 – 1:3. |  |  |  |  | Error? Not compatible with the previous requirement
6. Respiratory Rate. The number of breathing cycles every minute. |  | 'Respiratory Rate' is the number of breathing cycles every minute. | Comment
a. Range 10 – 30 breaths per minute in increments of 2 (only in mandatory mode) can be set by the user. | 29 | If the system is in 'mandatory mode' then the system shall enable the user to set the breathing rate in the range 10 to 30 breaths per minute | REQT | Mandatory | So 'mandatory mode' is a thing!
| 30 | If the system is in 'mandatory mode' then the system shall enable the user to set the breathing rate in increments of 2 breaths per minute | REQT | Mandatory
7. Tidal Volume (Vt) setting, if provided. The volume of gas flowing into the lungs during one inspiratory cycle |  | 'Tidal Volume' (Vt) is the volume of gas flowing into the lungs during one inspiratory cycle | Comment
a. Must have at least one setting of 400ml +/- 10 ml.# | 31 | The system shall have at least one Tidal volume setting of 400 ml +/- 10 ml | REQT | Mandatory
b. Ideally 350ml and 450 ml options. | 32 | The system shall have Tidal Volume setting of 350 ml +/- 10 ml | REQT | Optional
| 33 | The system shall have Tidal Volume setting of 450 ml +/- 10 ml | REQT | Optional
c. Optionally Range 250 – 600 ml in steps of 50ml. | 34 | The system shall enable the Tidal Volume to be set by a user in the range 250 to 600ml in steps of 50ml | REQT | Optional
d. Even more optionally up to 800 ml. | 35 | The system shall enable the Tidal Volume to be set by a user in the range 250 to 800ml in steps of 50ml | REQT | Optional
e. Optionally the ability to input body weight and have volume calculated as e.g. 6ml/kg of ideal body weight. | 36 | The system shall enable the calculation of Tidal Volume based on user-entered patient body-weight, using the formula: Tidal Volume = 6ml/Kg of Ideal body weight | REQT | Optional | Unclear: Does this mean to just do the calculation and display the result? Or does it mean that the Tidal Volume can be set by entering ideal body weight and have the system do the calculation before automatically setting the TV? Also unclear: the use of "E.g. 6ml/Kg" .. is this just one example of many potential calculations. Should this have been "I.e."?
Gas and electricity |  | Gas and electricity | Comment
1. Incoming Gas Supply. |  | 1. Incoming Gas Supply. | Comment
a. All gas connectors and hoses must comply with BS EN ISO 5359:2014+A1:2017, ISO 5359:2014/AMD 1:2017 and BS 2050: 1978 Electrical Conductivity. | 37 | All gas connectors must comply with BS EN ISO 5359:2014+A1:2017, ISO 5359:2014/AMD 1:2017 | REQT | Mandatory
| 38 | All gas hoses must comply with BS EN ISO 5359:2014+A1:2017, ISO 5359:2014/AMD 1:2017 | REQT | Mandatory
| 39 | All gas connectors and hoses must comply with BS 2050: 1978 Electrical Conductivity. | REQT | Mandatory
| 40 | All gas hoses must comply with BS 2050: 1978 Electrical Conductivity. | REQT | Mandatory
b. Must connect to wall pipeline oxygen supply via BS 5682:2015 compatible probes (Schrader). If hose not permanently fixed to machine, then must connect with NIST (Non-Interchangeable Screw Thread to ISO 18082:2014/AMD 1:2017). Oxygen pipeline pressure is approximately 3.7 – 4.5 bar. | 41 | The system shall connect to wall pipeline oxygen supply via BS 5682:2015 compatible probes (Schrader). | REQT | Mandatory
| 42 | If the system oxygen supply hose is not permanently fixed to machine, then it must connect with NIST (Non-Interchangeable Screw Thread to ISO 18082:2014/AMD 1:2017). | REQT | Mandatory
| 43 | The system shall be capable of operating with an Oygen supply input pressure in the range 3.7 – 4.5 bar. | REQT | Mandatory | Unclear: Changed from a comment to a requirement. Input pressure needs to be specified
c. Oxygen supply from wall outlets outside of ICU and theatres is limited to approximately 6-10 lpm (HTM_02-01_Part_A) and so provision for a gas reservoir will be required to manage peak inspiratory flow rates of up to 100 lpm |  | Oxygen supply from wall outlets outside of ICU and theatres is limited to approximately 6-10 lpm (HTM_02-01_Part_A) | Comment |  | Implicitely - this is not enough to support the required level of ventilation
| 44 | The system shall provide a gas reservoir to manage peak inspiratory flow rates of up to 100 lpm based on a minimum 6 lpm input | REQT | Mandatory | Incomplete : How long does the resevoir have to provide the 100 lpm for? Indefinately?
d. Average oxygen consumption must be no more than 3 lpm. This may be allowed to increase as greater certainty is gained over oxygen supply. | 45 | The system shall regulate the consumption of input Oxygen to below 3 Lpm | REQT | Mandatory
|  | The 3 lpm limit may be relaxed This may be allowed to increase as greater certainty is gained over oxygen supply. | Comment
e. Optionally can connect to wall pipeline Medical Air (MA4, NOT SA7) via BS 5682:2015 compatible probes. | 46 | The system shall connect to wall pipeline Medical Air (MA4, NOT SA7) via BS 5682:2015 compatible probes. | REQT | Optional
f. Optionally can connect to ISO 7396-2:2007 compatible Anaesthetic Gas Scavenging System or an external activated charcoal absorber (If inhaled anaesthetic agents are being used). | 47 | The system shall connect to an ISO 7396-2:2007 compatible Anaesthetic Gas Scavenging System | REQT | Optional | Ambigious because of 'OR' clause
| 48 | The system shall connect to an ISO 7396-2:2007 compatiblean external activated charcoal absorber | REQT | Optional | Ambigious because of 'OR' clause
g. Optionally can operate using an oxygen concentrator device for input oxygen, these will typically be limited to 10 lpm. | 49 | The system shall be capable of operating with an Oxygen supply from an Oxygen concentrator device | REQT | Optional
|  | Oxygen concentrator devices will typically be limited to less than 10 lmp | Comment
2. Electricity Supply |  | 2. Electricity Supply | Comment
a. Should connect to 240V mains. | 50 | The system shall operate from 240V AC mains supply | REQT | Mandatory
b. Should be PAT tested to the adapted IEC 60601, IEC 62353 standards | 51 | The system shall comply with the adapted IEC 60601 standard | REQT | Mandatory
| 52 | The system shall comply with the adapted IEC 62353 standard | REQT | Mandatory
c. Battery backup – see Unknown issues section. Must have 20 minutes back up battery in case of mains electricity failure. | 53 | The system shall be capable of operating without external power-supply for at least 20 minutes | REQT | Mandatory
d. Optionally hot swappable batteries so that it can be run on battery supply for an extended period, e.g. 2 hours for within hospital transfer. | 54 | The system shall be capable of running from 'hot swappable' batteries to enable it to run without an external power supply for periods up to 2 hours | REQT | Optional
e. Must avoid harmful RF or EM emissions that could interfere with other critical machinery. | 55 | The system shall not emit RF radiation that will interfere with other medical equipment | REQT | Mandatory | Unclear : Should refer to an EMC standard?
| 56 | The system shall not emit EM radiation that will interfere with other medical equipment | REQT | Mandatory | Not clear to me how 'RF' differs from 'EM' - other than EM is more general?
3. Gas supply to patient. |  | 3. Gas supply to patient. | Comment
a. User must be able to control inspired oxygen proportion (FiO2). The percentage of oxygen in the gas being breathed in by the patient. Room air is 21% oxygen. |  | Room air is 21% oxygen. | Comment
| 57 | (The system shall enable the user to control the Inspired Oxygen Proportion (Fi02) to 50%) OR (The system shall enable the user to control the Inspired Oxygen Proportion (Fi02) to 60%) | REQT | Mandatory
b. At least (50% or 60%) and 100% options | 58 | The system shall enable the user to control the Inspired Oxygen Proportion (Fi02) to 100% | REQT | Optional
c. Ideally variable between 30 – 100 % in 10% steps. | 59 | The system shall enable the user to control the Inspired Oxygen Proportion (Fi02) within the range 30% to 100% | REQT | Optional
| 60 | The system shall enable the user to control the Inspired Oxygen Proportion (Fi02) in 10% steps | REQT | Optional
d. Patient breathing system connections: the ventilator must present 22mm outside diameter (OD) ‘male’ standard connectors to ISO 5356-1:2015 on both outlet and inlet ports for connection to user supplied 22mm ‘female’ connectors on the breathing system. These must be rigid and robust (not plastic) and separated by a minimum of 10 cm between centres to accommodate filter HMEs. |  | Patient breathing system connections: | Comment
| 61 | The system shall present 22mm outside diameter (OD) ‘male’ standard connectors to ISO 5356-1:2015 on outlet ports for connection to user supplied 22mm ‘female’ connectors on the breathing system. | REQT | Mandatory
| 62 | The system shall present 22mm outside diameter (OD) ‘male’ standard connectors to ISO 5356-1:2015 on inlet ports for connection to user supplied 22mm ‘female’ connectors on the breathing system. | REQT | Mandatory
| 63 | The system's Patient breathing system connectors shall be ridged | REQT | Mandatory | Standard required? What does 'ridged' / 'Robust' mean in this context?
| 64 | The system's Patient breathing system connectors shall be robust | REQT | Mandatory
| 65 | The system's Patient breathing system connectors shall not be made of plastic | REQT | Mandatory
4. All elements in the gas pathway must meet biological safety and low- pressure oxygen safety standards, especially to minimise risk of fire or contamination of the patient’s airway. |  | Requirements [66] and [67] are intended to minimise risk of fire or contamination of the patient’s airway. | Comment
| 66 | All elements of the gas pathway within the System must meet biological safety standards | REQT | Mandatory | What are those standards?
| 67 | All elements of the gas pathway within the System must meet low-pressure oxygen safety standards | REQT | Mandatory | What are those standards?
Infection control
1. All parts coming into contact with the patient’s breath must be either disposable or able to be decontaminated between patients. | 68 | All parts of the System that come into contact with the patient’s breath must be either disposable or able to be decontaminated between patients. | REQT | Mandatory
2. All external surfaces must be cleanable in the likely event that they get respiratory secretions or blood splatter on them. Cleaning would be by healthcare workers manually wiping using an approved surface wipe with disinfectant or cloths and approved surface cleaning liquid. |  | Requirements [69] and [70] arrise since it is likely that the System shall get respiratory secretions or blood splatter on it | Comment
| 70 | All external surfaces of the System must be capable of being cleaned by wiping with an approved surface wipe with disinfectant | REQT | Mandatory | What are the 'approved surface wipes'? What are their properties? Implications for design?
| 71 | All external surfaces of the System must be capable of being cleaned by wiping with an approved cloths and approved surface cleaning liquids | REQT | Mandatory | What are the 'approved cleaning liquids'? What are their properties? Implications for design?
3. There will be a separately sourced HMEF-bacterial-viral filter between the machine and patient which may impact on resistance within the system, which may need to be accounted for with some designs. The pressure being delivered to the patient is the specified pressure. If the filter has a resistance of, say 2 cmH2O at 30 lpm, the ventilator needs to output 37 cmH2O to achieve a set 35 cmH2O at the patient. This will need further detailed consideration. Viral filtering filters may have much higher resistance that may be clinically relevant. |  | The pressure being delivered to the patient is the specified pressure. If the filter has a resistance of, say 2 cmH2O at 30 lpm, the ventilator needs to output 37 cmH2O to achieve a set 35 cmH2O at the patient. This will need further detailed consideration. Viral filtering filters may have much higher resistance that may be clinically relevant. | Comment |  | Note that requirement [72] may therefore be incomplete.
| 72 | The system shall be capable of meeting its pressure delivery requirements when fitted with an external HMEF-bacterial-viral filter between the System and the patient. | REQT | Mandatory
4. Optionally include facility for hot water humidifier to be included in breathing system. | 73 | The system shall include a hot water humidifer | REQT | Optional
Monitoring and Alarms |  | Monitoring and Alarms | Comment
1. Must alarm at: | 74 | The System shall generate an Alarm in the event of gas supply failure | REQT | Mandatory | Specific requirements for the 'Alarm' - Sounds? (How loud? Pitch? Intermittent? Silencable? At the machine or at a distance (central patient monitoring).
a. Gas or electricity supply failure. | 75 | The System shall generate an Alarm in the event of electrical supply failure | REQT | Mandatory | Specific requirements for the 'Alarm' - Sounds? (How loud? Pitch? Intermittent? Silencable?)
b. Machine switched off while in mandatory ventilation mode. | 76 | The System shall generate an Alarm in the event of it being switched off whilst in mandatory ventallation mode | REQT | Mandatory | Implies either a 'soft' on/off switch .. Or an indendependent alarm system that runs continually off of battery even if the main power to the device is switched off
c. Inspiratory airway pressure exceeded. | 77 | The System shall generate an Alarm in the event of [the maximum? ]Inspiratory airway pressure being exceeded. | REQT | Mandatory | This is an alarm associated with the limit setting of requirement [20] - correct?
d. Inspiratory and PEEP pressure not achieved (equivalent to disconnection alarm). | 78 | The System shall generate an Alarm in the event of Inspiratory airway pressure not being achieved | REQT | Mandatory | Original requirement ambigious because of 'AND'. Assume that either of these coonditions warrant an alarm
| 79 | The System shall generate an Alarm in the event of PEEP not being achieved | REQT | Mandatory
e. Tidal volume not achieved or exceeded. | 80 | The System shall generate an Alarm in the event that the Tidal volume exceeded | REQT | Mandatory | Original requirement ambigious because of 'AND'. Assume that either of these coonditions warrant an alarm
| 81 | The System shall generate an Alarm in the event that the Tidal volume is not achieved | REQT | Mandatory
| 82 | The System shall generate an Alarm if the backup power supply (battery) falls below [TBD assume 90%] charge levels | REQT | Mandatory | New requirement generated by Ventilator crowd - otherwise you have a dormant failure when on mains power but any connected battery is not capable of running the machine. Also need to warn the clinician of impedding 'hot swap' of battery or other mitigation action
2. Monitoring – the following should be continuously displayed so the user can verify.
a. Current settings of tidal volume, frequency, PEEP, FiO2, ventilation mode. | 83 | The system shall continuosly display the current settings for Tidal Volume | REQT | Mandatory
| 84 | The system shall continuosly display the current settings for Frequency | REQT | Mandatory | Frequency?
| 85 | The system shall continuosly display the current settings for PEEP | REQT | Mandatory
| 86 | The system shall continuosly display the current settings for FiO2 | REQT | Mandatory
| 87 | The system shall continuosly display the ventilation mode | REQT | Mandatory
b. Actual achieved rates of tidal volume, breathing rate, PEEP, plateau pressure, FiO2. | 88 | The system shall continuosly display achieved tidal volume | REQT | Mandatory
| 89 | The system shall continuosly display achieved breathing rate | REQT | Mandatory | Measured over what period? How many breaths (or amount of time) should this average to count rate?
| 90 | The system shall continuosly display achieved PEEP | REQT | Mandatory
| 91 | The system shall continuosly display achieved Plateau pressure | REQT | Mandatory
| 92 | The system shall continuosly display achieved FiO2 | REQT | Mandatory
c. If it exists, in pressure support mode there must be real time confirmation of each patient breath and an alarm if below acceptable range. | 93 | If the system includes a 'Pressure Support Mode' then the system shall provide a real-time confirmation of each patient breath | REQT | Mandatory | What type of confirmation? A 'beep'? A visual indication? Both?
| 94 | If the system includes a 'Pressure Support Mode' then the system shall generate an alarm if the patients breathing rate drops below an accepted range | REQT | Mandatory | Unclear: (1) Is this about breathing rate or breathing volume? (2) What is the 'accepted range' ? .. As defined in previous requirements?
d. Optionally CO2 monitoring included. | 95 | The System shall provide CO2 monitoring | REQT | Optional | Incomplete: Monitoring of CO2 in patients exhaled breath? Or is this CO2 on the input to the patient? Or within the system? Or ambient?
Miscellaneous |  | Miscellaneous | Comment
1. Must be reliable. It must have 100% duty cycle for up to 14 days. | 96 | The system shall have an intrinsic Mean Time Betweek Failures (MTBFi) of [ TBD MTBFi number needed ] | REQT | Mandatory | Very unclear. What is "up to 14 day". 1 failure in 14 days (336 hours)? .. But note .. This is not an MTBF but a minimum requirement
2. Optionally it can be used beyond 14 days, the expected durability must be specified. | 97 | The system shall be durable (capable of operating) continuosly for a period in excess of 336 hours - supllier to state achieved durability | REQT | Optional | Reliability' has now become 'Durability' which is intended here? Is this an intrinsic reliability requirement or a wear-out (systematic) failure rate requirement?
3. Can be floor standing. | 98 | The system shall be floor standing | REQT | Optional | Not clear if this is a requirement or a comment
4. Ideally small and light enough to mount on patient bed and orientation independent functioning. | 99 | The system shal fit within a volume [Height TBD ] by [ Width TBD ] by [Depth TBD ] | REQT | Optional
| 100 | The System shall weigh less than [Mass Kg TBD] | REQT | Optional
| 101 | The System shall be capable of operating independent or orientation | REQT | Optional | Unclear: Does this 'mechanically; (i.e. running upsidedown) or with a display that can be oriented to enable user operation whilst the device is fitted to a patient bed
5. Should be as robust as possible. For example, it may be dropped from bed height to floor. | 102 | The System shall be capable of sustaining a drop from 1m onto a hard floor | REQT | Optional
6. It must be intuitive to use for qualified medical personnel, but these may not be specialists in ventilator use. |  | It must be intuitive to use for qualified medical personnel, but these may not be specialists in ventilator use. | Comment |  | Intuitive' is not translatable into a requirement. Following items adopted as requirement. This item could be translated into a work requirement for usability testing.
a. Must not require more than 30 minutes training for a doctor with some experience of ventilator use. | 103 | The system shall enable a user [ with some experience of ventilator use ] no more than 30 minutes training to operate it | REQT | Mandatory | Does this have to be a 'doctor'? Our reading is that these are usually used by a specialist nurse. Changed 'Doctor' to 'user'
b. Must include Instructions for Use. | 104 | The System shall include instructions for its use. | REQT | Mandatory
c. Ideally instructions for use should be built into the labelling of the ventilator, e.g. with ‘connect this to wall’ etc. | 105 | The System shall include user-visible labelling for all external connections | REQT | Mandatory
d. Must include clear labelling of all critical functions and controls using standard terms, pictograms and colours that will be readily recognised by UK healthcare staff. | 106 | The System shall include user-visible labelling providing a succinct set of instructions for use | REQT | Mandatory
| 107 | The System shall include user-visible labelling of all critical controls | REQT | Mandatory | What is a 'critical control'? Whch controls are critical?
| 108 | The System shall include user-visible labelling of all critical functions | REQT | Mandatory | Unclear: How can a 'function' be labelled? Does this mean 'settings'? Or labelling of what the dispays mean
| 109 | All labelling withing the System shall use standard terms | REQT | Mandatory | Where are the 'standard terms' defined? Or do these need to go through an approval / testing process?
| 110 | All labelling withing the System shall use standard pictograms that will be readily recognised by UK healthcare staff. | REQT | Mandatory | Where are the 'standard pictograms' defined?
| 111 | All labelling withing the System shall use standard colours that will be readily recognised by UK healthcare staff. | REQT | Mandatory | What are the 'standard colours'? Where are these defined?
7. Must have transparent design, supply chain, manufacture, quality assurance and testing processes that are of sufficient quality to enable MHRA officials to deem appropriate for usage in exceptional circumstances. | 112 | The supplier must have transparent design proceses | SOW | Mandatory
| 113 | The supplier must have transparent supply-chain proceses | SOW | Mandatory
| 114 | The supplier must have transparent manufacturing proceses | SOW | Mandatory
| 115 | The supplier must have transparent quality assurance proceses | SOW | Mandatory
| 116 | The supplier must have transparent testing proceses | SOW | Mandatory
8. Must not be excessively cumbersome so that it would impede hospital operations or prevent easy movement within hospital premises. | 117 | The System shall not impede hospital operations | REQT | REQT | Unclear: This sounds like a size / weight requirement. Needs to be further developed
| 118 | The System shall not impeded movement within hospital premises | REQT | REQT | Unclear: This sounds like a size / weight requirement. Needs to be further developed
9. Must be made from materials and parts readily available in the UK supply chain (anticipating increasing global restrictions on freight movement). | 119 | The System shall be constructed of materials that are readily available in the UK supply chain | REQT | REQT
| 120 | The System shall be constructed of parts that are readily available in the UK supply chain | REQT | REQT
10. Standards – there are many standards that exist in this area. Below is a list of the most relevant ones. They are not formal regulatory requirements, but many are harmonised against regulatory requirements. Consider them as helpful advisory standards for now. MHRA will lead an exercise to define which can be ‘safely’ relaxed for this emergency situation. |  | 10. Standards – there are many standards that exist in this area. Below is a list of the most relevant ones. They are not formal regulatory requirements, but many are harmonised against regulatory requirements. Consider them as helpful advisory standards for now. MHRA will lead an exercise to define which can be ‘safely’ relaxed for this emergency situation. | Comment
a. BS EN 794-3:1998 +A2:2009 Particular requirements for emergency and transport ventilators |  | BS EN 794-3:1998 +A2:2009 Particular requirements for emergency and transport ventilators | Comment
b. ISO 10651-3:1997 Lung Ventilators for Medical Use - Emergency and Transport |  | ISO 10651-3:1997 Lung Ventilators for Medical Use - Emergency and Transport | Comment
c. BS ISO 80601-2-84:2018 Medical electrical equipment. Part 2-84. Particular requirements for basic safety and essential performance of emergency and transport ventilators – especially the parts on ‘patient gas pathway’ safety (very similar to IEC 60601) |  | BS ISO 80601-2-84:2018 Medical electrical equipment. Part 2-84. Particular requirements for basic safety and essential performance of emergency and transport ventilators – especially the parts on ‘patient gas pathway’ safety (very similar to IEC 60601) | Comment
d. ISO 80601-2-12:2020 Medical electrical equipment — Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators |  | ISO 80601-2-12:2020 Medical electrical equipment — Part 2-12: Particular requirements for basic safety and essential performance of critical care ventilators | Comment
e. BS ISO 19223:2019 Lung ventilators and related equipment. Vocabulary and semantics |  | BS ISO 19223:2019 Lung ventilators and related equipment. Vocabulary and semantics | Comment
Testing |  | Testing | Comment
1. It is accepted that full demonstration of compliance to ISO 80601-2-12:2020 is unrealistic in the time frame required for development. Nevertheless, compliance with the essential safety standards must be demonstrated for patient safety. |  | It is accepted that full demonstration of compliance to ISO 80601-2-12:2020 is unrealistic in the time frame required for development.
| 121 | The Supplier shall demonstrate the System's compliance with the essential safety standards for patient safety. | SOW | Mandatory | Incomplete: What are the "essential safety standards"?
2. It is not anticipated that devices will be CE marked and approval by the MHRA will be through the “Exceptional use of non-CE marked medical devices” route. |  | It is not anticipated that devices will be CE marked and approval by the MHRA will be through the “Exceptional use of non-CE marked medical devices” route. | Comment |  | Completely unclear .. Missing words?
3. When the current emergency has passed these devices will NOT be usable for routine care unless they have been CE marked through the Medical Device Regulations. The device must display a prominent indelible label to this effect. |  | When the current emergency has passed these devices will NOT be usable for routine care unless they have been CE marked through the Medical Device Regulations. | Comment
| 122 | The System shall display a prominent indelible label identifying that it is not suitable for routine care after the current emergency has past | REQT | Mandatory
4. Usability testing at both prototype and final production stages will be required. This should be done as a short Formative Usability Test to ISO62366 (this can be done in a day) in a realistic environment if possible. The user will be wearing complex protective clothing which includes: Eye goggles (in addition to spectacles if worn), Face shield, Plastic apron, Surgical gown, Two layers of gloves, usually nitrile non- handed small, medium, large variants, Gloves are donned in layers and sticky taped onto sleeves of gown in between layers | 123 | The supplier shall complete usability testing of any prototype device | SOW | Mandatory | Assume that this means 'Type testing' rather than 'Item testing'
| 124 | The supplier shall complete usability testing of any final production device | SOW | Mandatory | Assume that this means 'Type testing' rather than 'Item testing'
| 125 | Usability testing shall be conducted as a short Formative Usability Test to ISO62366 | SOW | Mandatory
| 126 | Usability Testing shall be completed in a realistic environment. | SOW | Optional
|  | The user will be wearing complex protective clothing which includes: Eye goggles (in addition to spectacles if worn), Face shield, Plastic apron, Surgical gown, Two layers of gloves, usually nitrile non- handed small, medium, large variants, Gloves are donned in layers and sticky taped onto sleeves of gown in between layers | Comment |  | There is a whole set of 'hidden' ergonomics / usability requirements here. These should be factored out into a separate usabiloty / ergonmics requirements section
5. The user must be able to instantly see the settings selected and be able to easily operate all controls while dressed in protective gear. They may be required to remain so clothed and operating the ventilator for a number of hours without breaks. |  | Users may be required to remain so clothed and operating the ventilator for a number of hours without breaks. | Comment
| 127 | The System shall enable users to instantly see machine settings whilst dressed in protective equipment | REQT | Mandatory
| 128 | The System shall enable users to [easily] operate all controls whilst dressed in protective equipment | REQT | Mandatory | Not clear what 'easily' means? Usability testing? Standards?
Unknown Issues |  | Unknown Issues | Comment
1. How plentiful is 4 bar oxygen supply? |  | 1. How plentiful is 4 bar oxygen supply? | Comment
a. Absolute minimum oxygen requirement is the human consumption of about 250 ml/min in a healthy person but upto 500 ml/min in severe sepsis. However, achieving this is only possible if certain breathing system designs are used and ‘driving’ gas is done by air. |  | a. Absolute minimum oxygen requirement is the human consumption of about 250 ml/min in a healthy person but upto 500 ml/min in severe sepsis. However, achieving this is only possible if certain breathing system designs are used and ‘driving’ gas is done by air. | Comment
i. Specifically, would have to use circle breathing system with active CO2 absorption. Is sufficient soda lime available? |  | i. Specifically, would have to use circle breathing system with active CO2 absorption. Is sufficient soda lime available? | Comment
b. If consumption in the range 1-2 l/min is acceptable then a wider range of designs is possible, but some very basic designs are not. |  | b. If consumption in the range 1-2 l/min is acceptable then a wider range of designs is possible, but some very basic designs are not. | Comment
c. If consumption in the range 10l/min is acceptable then any possible design can be considered. |  | c. If consumption in the range 10l/min is acceptable then any possible design can be considered. | Comment
2. What is the resistance of HMEF-bacterial-viral filters that are to be used with the ventilator? Is it clinically relevant? |  | 2. What is the resistance of HMEF-bacterial-viral filters that are to be used with the ventilator? Is it clinically relevant? | Comment
3. Is there any need to consider running from only low-pressure oxygen e.g. from a concentrator? This makes design more complex. |  | 3. Is there any need to consider running from only low-pressure oxygen e.g. from a concentrator? This makes design more complex. | Comment
4. How plentiful is the supply of syringe drivers and drugs for sedation? |  | 4. How plentiful is the supply of syringe drivers and drugs for sedation? | Comment
a. If limited, then a vaporiser could be used to vaporise Isoflurane for sedation. |  | a. If limited, then a vaporiser could be used to vaporise Isoflurane for sedation. | Comment
b. This would need certain breathing system designs, mandatory AGSS and a supply of vaporisers. |  | b. This would need certain breathing system designs, mandatory AGSS and a supply of vaporisers. | Comment
5. If monitoring can be done by another machine it could be left out of the ventilator, but essential parameters must be available to the clinician. |  | 5. If monitoring can be done by another machine it could be left out of the ventilator, but essential parameters must be available to the clinician. | Comment
6. Battery backup – Every current ventilator used inside hospitals has a battery backup, so users will expect it to be there and will behave as if it is, e.g. unplug it from the wall in order to rearrange cables or while manoeuvring the patient. However, this needs very careful thought to balance the risks. Including this in the spec means instantly trying to source 30,000 large, heavy batteries. Specifying a DC voltage (ie 12VDC) may well be the most sensible for the machine working voltage. Need the advice of an electronic engineer with military/resource limited experience before specifying anything here. It needs to be got right first time. |  | 6. Battery backup – Every current ventilator used inside hospitals has a battery backup, so users will expect it to be there and will behave as if it is, e.g. unplug it from the wall in order to rearrange cables or while manoeuvring the patient. However, this needs very careful thought to balance the risks. Including this in the spec means instantly trying to source 30,000 large, heavy batteries. Specifying a DC voltage (ie 12VDC) may well be the most sensible for the machine working voltage. Need the advice of an electronic engineer with military/resource limited experience before specifying anything here. It needs to be got right first time. | Comment
Glossary |  | Glossary | Comment
ARDS – Acute Respiratory Distress Syndrome: a life-threatening form of respiratory failure where the lungs become severely inflamed due to an infection or injury and can't provide the body's vital organs with enough oxygen. |  | ARDS – Acute Respiratory Distress Syndrome: a life-threatening form of respiratory failure where the lungs become severely inflamed due to an infection or injury and can't provide the body's vital organs with enough oxygen. | Comment
SIMV-PC – Synchronized Intermittent Mandatory Ventilation – Pressure Controlled: a mode of ventilation where the patient is allowed to take spontaneous breaths, the machine will assist the patients breathing when a spontaneous breath is taken. If the patient does not make a pre-set number of breaths a minute (i.e. 10) the machine provides mechanical ventilation to provide the set number. |  | SIMV-PC – Synchronized Intermittent Mandatory Ventilation – Pressure Controlled: a mode of ventilation where the patient is allowed to take spontaneous breaths, the machine will assist the patients breathing when a spontaneous breath is taken. If the patient does not make a pre-set number of breaths a minute (i.e. 10) the machine provides mechanical ventilation to provide the set number. | Comment
PRVC – Pressure Regulated Volume Controlled: A mode of ventilation where a set tidal volume is delivered to the patient while maintain the lowest pressure possible in the airway, to avoid trauma. |  | PRVC – Pressure Regulated Volume Controlled: A mode of ventilation where a set tidal volume is delivered to the patient while maintain the lowest pressure possible in the airway, to avoid trauma. | Comment
BIPAP – Bilevel Positive Airway Pressure: a non-invasive ventilation mode that provides different levels of pressure when the patient inhales and exhales. |  | BIPAP – Bilevel Positive Airway Pressure: a non-invasive ventilation mode that provides different levels of pressure when the patient inhales and exhales. | Comment
PEEP – Positive End-Expiratory Pressure: The lower pressure applied to the patient’s airway to allow them to breathe out, but not too much. |  | PEEP – Positive End-Expiratory Pressure: The lower pressure applied to the patient’s airway to allow them to breathe out, but not too much. | Comment
EPAP – Expiatory Positive Airway Pressure: Similar to PEEP, pressure applied to the airway on patient expiration to prevent collapse of the airway. |  | EPAP – Expiatory Positive Airway Pressure: Similar to PEEP, pressure applied to the airway on patient expiration to prevent collapse of the airway. | Comment
HMEF – Heat and Moisture Exchange Filter: device fitted to the patient end of the breathing system, contains hydrophobic medium that absorbs heat and moisture from the patients exhaled breath and uses absorbed moisture to humidify inhaled gases. Can also filter bacteria and viruses, this will be used on all patients. WARNING can affect delivered pressure. |  | HMEF – Heat and Moisture Exchange Filter: device fitted to the patient end of the breathing system, contains hydrophobic medium that absorbs heat and moisture from the patients exhaled breath and uses absorbed moisture to humidify inhaled gases. Can also filter bacteria and viruses, this will be used on all patients. WARNING can affect delivered pressure. | Comment
RF – Radio Frequency: Many medical devices are sensitive to RF interference. Care should be taken to ensure that this is kept to a minimum. |  | RF – Radio Frequency: Many medical devices are sensitive to RF interference. Care should be taken to ensure that this is kept to a minimum. | Comment
EM – Electro Magnetic Emissions: Many medical devices are sensitive to EM interference. Care should be taken to ensure that this is kept to a minimum. |  | EM – Electro Magnetic Emissions: Many medical devices are sensitive to EM interference. Care should be taken to ensure that this is kept to a minimum. | Comment
Fi02 – Fraction of inspired oxygen: concentration of oxygen in the gas mixture that the patient inhales |  | Fi02 – Fraction of inspired oxygen: concentration of oxygen in the gas mixture that the patient inhales | Comment
AGSS – Anaesthetic gas scavenging system: where anaesthetic agents have been included in the gas mixture, this system is used to collect and remove exhales gas to avoid exposure to health care professionals. |  | AGSS – Anaesthetic gas scavenging system: where anaesthetic agents have been included in the gas mixture, this system is used to collect and remove exhales gas to avoid exposure to health care professionals. | Comment